## Slide 1: Diagnostic Center
Annual QA Analysis 2025

## Slide 2: Executive Summary
OVERALL STATISTICS: â€¢ Total Specimens: 324,119 â€¢ Total Rejected: 118 (0.036%) â€¢ Average QC Acceptance: 93.1% CRITICAL ISSUES: 1. Hematology - April data corruption 2. Routine Chemistry - QC dropped to 82.3% (Dec) 3. Special Chemistry - Below 95% target 4. Molecular Pathology & Histopathology - No QC performed TOTAL PANIC NOT INFORMED: 16 cases

## Slide 3: ðŸš¨ CRITICAL ALERTS
IMMEDIATE ACTION REQUIRED: 1. HEMATOLOGY - APRIL DATA CORRUPTION â€¢ 190 failed QC out of 0 tests run (Impossible) â€¢ Shows 0% QC acceptance - Data entry error â€¢ 1500% panic notification rate (Impossible) 2. ROUTINE CHEMISTRY - CRITICAL DECLINE â€¢ December QC dropped to 82.3% â€¢ 5 consecutive months below 90% â€¢ Equipment calibration needed 3. MISSED PANIC NOTIFICATIONS â€¢ 16 critical results not informed â€¢ Patient safety risk 4. NO QC IN MOLECULAR & HISTOPATHOLOGY

## Slide 4: Hematology Analysis
STATISTICS: â€¢ Total Specimens: 111,307 â€¢ Rejected Samples: 80 â€¢ QC Acceptance: 90.9% â€¢ Panic Notification: 98.3% MONTHLY HIGHLIGHTS: â€¢ Peak Month: October (12,936 specimens) â€¢ Low QC: December (96.7%) â€¢ April: DATA ERROR - Results unreliable ISSUES: âš April data corrupted - investigate âš 16 panic results not informed âš QC below 95% target

## Slide 5: Routine Chemistry - CONCERN
STATISTICS: â€¢ Total Specimens: 106,357 â€¢ QC Acceptance: 88.8% (AVG) â€¢ Panic Notification: 100% â€¢ December QC: 82.3% (LOWEST) QC TREND: â€¢ January: 92% â†’ February: 84% â€¢ March: 95% â†’ April: 87% â€¢ Continues declining... â€¢ December: 82.3% âš ï¸ CONCERNS: ðŸš¨ NEVER reached 95% target ðŸš¨ Equipment calibration needed ðŸš¨ Declining trend continues

## Slide 6: Special Chemistry
STATISTICS: â€¢ Total Specimens: 60,944 â€¢ QC Acceptance: 90.5% (AVG) â€¢ Panic Notification: 100% â€¢ Lowest: February (81%) MONTHLY QC: â€¢ Jan: 89% | Feb: 81% | Mar: 93% â€¢ Apr: 83% | May: 85% | Jun: 89.5% â€¢ Jul: 95.5% | Aug: 94.5% | Sep: 96.3% â€¢ Oct: 95% | Nov: 96.3% | Dec: 93.5% STATUS: âœ“ Improving in recent months âš Below 95% for most of year âš External QA documentation needed

## Slide 7: Microbiology
STATISTICS: â€¢ Total Specimens: 42,882 â€¢ QC Acceptance: 99.7% (BEST) â€¢ Rejected Samples: 5 ISSUES: âš October: 0% QC acceptance âš No External QA documented âš Only 1 month with QC failure RECOMMENDATIONS: â€¢ Maintain good QC performance â€¢ Document External QA â€¢ Investigate October anomaly

## Slide 8: Molecular & Histopathology
âš ï¸ NO QC PERFORMED IN BOTH SECTIONS MOLECULAR PATHOLOGY: â€¢ Specimens: 620 â€¢ QC: 0% â€¢ External QA: 0% HISTOPATHOLOGY: â€¢ Specimens: 2,009 â€¢ QC: 0% â€¢ External QA: 0% ACTION REQUIRED: ðŸš¨ Implement QC program ðŸš¨ Start External QA ðŸš¨ Monitor outsourcing quality

## Slide 9: Recommendations
IMMEDIATE (This Week): 1. Fix April Hematology data 2. Investigate Routine Chemistry equipment 3. Increase QC frequency HIGH PRIORITY (This Month): 1. Implement automated panic alerts 2. Resume External QA for all sections 3. Start QC in Molecular/Histopathology 4. Equipment maintenance schedule ONGOING MONITORING: 1. Track monthly QC (target: 95%+) 2. Data integrity audits 3. Staffing for high-volume periods